## Slide 1: ASTM F748 Selecting Biological Test Methods
Kenneth R. St. John, PhD
Chairman
F04.16 Subcommittee on Biocompatibility Test Methods

## Slide 2: Conflict of Interest Statement
Prior research grants:
  Implex Corporation
  Depuy Synthes
  Globus Medical
Consulting
  Depuy Synthes
  Globus Medical
  Zimmer Biomet
Nothing of value has been received concerning the topics covered, other than travel support for participating in this meeting from the Standards Alliance and ASTM International

## Slide 3
Recommendations for test methods to be used to determine/verify biocompatibility of materials used in the construction of medical devices based upon:
  Intended application (Tissue contacted)
  Length of tissue contact
Makes reference to ASTM test methods and other relevant documents
Actual devices may be tested using the same principles, under certain circumstances

## Slide 4: Classification
Provide means of modifying selection of tests
  Time of contact
  Type of contact
Time of Contact
  Intraoperative – Less than 24 hours
  Short-term – Up to an including 30 days
  Chronic – Greater than 30 days

## Slide 5: Classification
Type of Contact
  External Devices
    Contacting Intact Body Surfaces
    Contacting Breached Body Surfaces
  Externally Communicating Devices
    Intact Natural Channels
    Body Tissues and Fluids
    Blood Path, Indirect
    Blood Path, Direct
  Implanted Long-Term Devices
    Contacting Bones
    Residing in the Subcutaneous Space
    Contacting Soft Tissue and Tissue Fluids
    Contacting Blood

## Slide 6: Examples Given for Categories

## Slide 7: F748 Test Selection Matrix

## Slide 8: Advice Given for Test Selection

## Slide 9: Many Tests Use Extracts
Different extraction fluids
  Saline
  Vegetable oils
  Others
    Polyethylene glycol
    DMSO
Extraction times and temperatures
Choose the highest temperature that will not degrade or deform the materials

## Slide 10: Cell Culture Test Methods

## Slide 11: Specialized Cell Culture Tests

## Slide 12: Scoring Guidance (F895)

## Slide 13: Control Material Guidance (F813)

## Slide 14: Control Material Guidance
For all types of tests
“Negative” control is usually a material that has been generally recognized as safe with extensive use
Polymers
  USP negative control plastic (polyethylene)
  UHMWPE (F648)
Metals
  316L stainless steel (F138)
  Titanium (Commercially Pure or alloy) (F67, F136)
  Cobalt Alloy (F75, F1537)
Ceramics
  Alumina (F603)
  Zirconia (ISO 13356, ASTM F1873 withdrawn)

## Slide 15: Specimen Creation
Entire sample or device may occasionally be used.
Usually a smaller sample is cut from supplied material sample or device
  Limitations of cell culture well
  Maximize surface area in contact with cells
  An extract may be used (F619) and may be placed on a filter disk for testing

## Slide 16: Results Reporting
Test method standards detail the information to be reported
Generally, “Pass” or “Fail” is not specified
  Response as compared to “negative” control
  Response reported as a score for cell culture testing
  Not all users are satisfied with this

## Slide 17: In Vivo Testing - Sensitization

## Slide 18: In Vivo Testing - Irritation
USP Irritation Test - Intracutaneous Injection Test
Mucous Membrane Irritation Test – No ASTM Standard

## Slide 19: In Vivo Testing – Short Term

## Slide 20: In Vivo Testing – Short Term
Provides for implantation for 7 days and 30 days – allows for 90 days as an option
May include sample or actual device
Gross appearance of tissues at implant site
Histopathology of implant site tissues
Rabbits – three animals at each time period
  Four test and four control in each rabbit

## Slide 21: In Vivo Testing – Long Term
Provide for implantation for a year or more
May include sample or actual device
Complete necropsy performed
  Systemic effects
  Local effects
  Lymph nodes
  Fluids (Blood etc.)
  Filtration/detox organs
    Liver
    Kidneys
    Brain – Arterial Circulation
    Lungs – Venous Circulation

## Slide 22: In Vivo Testing – Long Term

## Slide 23: F981 – Long Term Implantation
Standard specimen is cylindrical
  Range of sizes, depending upon site
  Test and control samples in same animal
    What if there is a systemic effect?
    How about absorbable materials? (F1983 in following slide)
  Should reflect final processing
    Clean, defect free, sterile, final packaging
Sites
  Muscle
  Bone
  Other sites?
Concern for porous materials – should be the same characteristics as will be used

## Slide 24: F981 – Long Term Implantation
Several models
  Rat (one test, one control)
  Rabbit (four test and two control)
  Larger animals (eight test and four controls)
  Rabbits or larger at least 16 test and 8 control per experimental time period (possibly fewer in rats)
Time periods
  12 weeks (usually, unless done in F763)
  26 weeks
  52 weeks
  Sometimes 104 weeks? (not in standard)

## Slide 25: Suggested Scoring – F981

## Slide 26: Absorbable Materials
Evaluate toxicity of degradation products first
Implantation times based upon degradation rates (50% and 100% mass loss)
Evaluation similar to F981
  Should controls and tests be in same animal?
  Evaluation of target organs not required, but are collected and preserved with strong encouragement for analysis

## Slide 27: Carcinogenicity
Test for lifetime of animals
  Animals naturally get tumors
  Compare frequency and types of tumors
Rarely performed for medical materials
Mice not recommended. Just rats
Complete necropsy at termination
Based on OECD protocol for chemicals

## Slide 28: Genotoxicity
Other tests
  Ames test
  Many tests are reverse mutation studies
    Look for reversal of a mutation
    Easier to assess
  Look for other sources of tests

## Slide 29: Genotoxicity
E1263 Guide for Conduct of Micronucleus Assays in Mammalian Bone Marrow Erythrocytes - Discontinued
E1280 Guide for Performing the Mouse Lymphoma Assay for Mammalian Cell Mutagenicity - Discontinued
E1397 Practice for In Vitro Rat Hepatocyte DNA Repair Assay - Discontinued
E1398 Practice for In Vivo Rat Hepatocyte DNA Repair Assay - Discontinued
E2186 Guide for Determining DNA Single-Strand Damage in Eukaryotic Cells Using the Comet Assay – Different ASTM Committee

## Slide 30: Particles

## Slide 31: Immunotoxicity? - Really Blood Contact
In addition to sensitization tests

## Slide 32: Blood Compatibility
All in vitro tests – ?Collection and storage?

## Slide 33: Neurotoxicity
Supplement to F748
Address additional issues with possible neurotoxicity testing
(More discussion in later talk)