## Slide 1: Case Study A New Motion-Sparing Spinal Disc Replacement
Kenneth R. St. John, PhD
Chairman
F04.16 Subcommittee on Biocompatibility Test Methods

## Slide 2: Components
Cobalt/Chromium alloy spherical (portion of a sphere) plate – Non nickel containing
Carbon fiber reinforced PEEK concave receiving cup attached by screws to a Ti-6Al-4V plate with a porous CP Titanium bone contact surface
Plasma-sprayed CP titanium bone contact surface on Co/Cr plate
Sheath to encapsulate particles?

## Slide 3: Concept
Co/Cr use as a bearing surface component has a long history but recent international controversies about metal-on-metal bearings may be a concern
Use of titanium and titanium alloys may improve osseointegration
Will a galvanic couple be formed between cobalt alloy and titanium alloy?
  Normally not a problem, but should be considered and perhaps tested

## Slide 4: Concept
While PEEK has seen use in spinal surgery, its use as a bearing surface is new, although it has been researched for hip prostheses
PEEK may have a greater wear resistance than UHMWPE
Will carbon fiber reinforcement of PEEK improve wear resistance or other mechanical properties?

## Slide 5: Concept
Will a polyurethane sheath around the bearing surface improve safety?
  Capture particles
  How about fatigue of PU?
  This decision may well be a trade-off in the design decision

## Slide 6: F748 Category

## Slide 7: Recommendations from F748
Cell culture cytotoxicity
Sensitization
Skin irritation or intracutaneous
Systemic toxicity
Pyrogen test
Short term implantation
Long term implantation
Immune response
Genotoxicity
Carcinogenicity

## Slide 8: Neurotoxicity
Supplement to F748
Address additional issues with possible neurotoxicity testing

## Slide 9: F2901 Neurotoxicity
Clinically relevant implantation studies
  Including neurobehavioral studies
  Recommended specific histological tests for neurodegeneration
Wear particle testing
  Proximity to spinal cord
Developmental neurotoxicity
Journal article references for testing that have been previously conducted

## Slide 10: Before Testing Selection
Determine whether testing information is available from literature or vendor for the materials being used
  Published studies
  Vendor biocompatibility testing
    Provide to manufacturer?
    File with regulators on a confidential basis for reference by manufacturer?
Do materials have long history of safe use? – May mean that biocompatibility can be inferred or testing requirements reduced – if can prove equivalent
  Polymers – same vendor?

## Slide 11: Cell Culture Cytotoxicity
Selection of test method
  F813 – Direct Contact
  F895 – Agar Overlay
  F1903 – Response to Particles
Substitute a neural cell line (if available)?
Use extract or material coupon?
  Extractables more appropriate for polymer than for metals
Which materials need to be tested?

## Slide 12: Sensitization
Is there any reason to expect sensitization?
  Materials with potential to elicit immune response?
  Perhaps should be performed to remove any question – opinions differ
    Many testing organizations advocate it regardless of application
Should patients be tested for sensitivity before implantation of devices?
  Rare done but advocated by some experts

## Slide 13: Skin Irritation or Intracutaneous
F719 – Primary Skin Irritation (Rabbit)
  Abraded and intact skin
  Direct contact for 24 hours
  Assess at removal and 24 and 48 hours later
F749 – Intracutaneous Injection (Rabbit)
  Intracutaneous is injection of an extract – does an extract have any meaning to this application?
Not being implanted in contact with skin but tests might be meaningful
  Contact with soft tissues in the spine (muscle & dura?

## Slide 14: Systemic Toxicity
F750 – Systemic Injection in the Mouse
  Extract
As with previously mentioned methods, when an extract is used, may not be testing all opportunities for problems
  Extracts evolved from testing of polymers
  Probably should assay extracts for content
  Metals and ceramics (may) have processing aids – but not monomers, plasticizers, catalysts, etc
Implantation studies may end up being more important

## Slide 15: Pyrogen Test
Should probably be performed on finished product to look for possible processing issues
LAL in vitro test is supplanting rabbit pyrogen test for bacterial endotoxins
Material pyrogenicity not measured by LAL
May be a process validation test
Major US class action lawsuit related to bacterial endotoxins settled for in excess of $1B

## Slide 16: Short Term Implantation
Screen for early adverse tissue reactions that might mean that material selection should be reconsidered
  Less possibility of unsuccessful long term study
  Reduce time and cost of learning of a problem
Standardized implantation test – F763
  Bone or muscle
Subcutaneous – F1408 – Not really appropriate for this device

## Slide 17: Long Term Implantation
F981 – Muscle or bone – suggested sites listed
  Site specific sites could also be used
    Contact with dura
    Intra- or inter-vertebral?
  Uses test samples or coupons. Not normally actual devices
    Should particulates also be tested, rather than the use of F1904?
Customized testing combining actual devices with long term implantation?

## Slide 18: Long Term Implantation
Use of actual (or sized down) devices in experimental animals may be an option
Humans are the only bipedal animals – loads and loading directions will be different
Many (most?) animals have much higher intervertebral range of motion than humans – may lead to damage or failures that might not occur in human use.

## Slide 19: Immune Response
Do we really need to be concerned about immune response?
None of the materials contain nickel
This testing may be expected for completeness but may have little relevance
Animal immune systems may differ from human – contact history, etc.
If a patient is sensitized to a material, testing in any other model will not address

## Slide 20: Genotoxicity
In a majority of cases (including this one), there is no reason to consider genotoxicity testing if materials have seen use previously
Genotoxicity could possibly be required for a truly new material

## Slide 21: Carcinogenicity
Carcinogenicity testing is very rarely done – test methods are more appropriate for chemicals than solid materials
An unpublished study was done on carbon fibers – no increase in tumor type, site, or incidence over sham operated

## Slide 22: Wear and Particulate
Note: Referencing this paper is only an example and implies no judgement about the device mentioned

## Slide 23: Wear and Particulate
Regulatory agencies appear to be sensitive to wear particles that may be generated by spinal prostheses
F1903 and F1904 offer in vitro and in vivo methods for assessing tissue response to particulate
Laboratory wear testing and clinical trial retrievals may yield particle size and mass information.
F561 and F1877 offer retrieval and characterization recommendations

## Slide 24: Wear and Particulate
Retrievals are showing that particulate is being generated by load bearing spinal devices
Research and clinical results have shown that wear particles from hip and knee joint prostheses can lead to destruction of surrounding tissues due to cytokine release from phagocytic cells
What testing is needed to determine whether particles that may be released may cause damage to bone, soft tissue and neural tissues?

## Slide 25: Wear and Particulate
Decision to be made for this device
  Add elastomeric envelope to contain particulate?
  How would lubrication occur?
  What is the risk of adding an additional material to design?
  Will elastomer fatigue with repeated flexing?

## Slide 26: Conclusions
All laboratory and animal testing appear to suggest safety
Perform laboratory simulation testing to look for potential problems
Example
  Make sure lab simulation mimics actual loading and usage
Determine whether results of mechanical testing suggest other biocompatibility testing
  Wear?

## Slide 27: Conclusions
Proceed to clinical trials
  Small numbers of patients?
  Full cohort?
Any retrievals from clinical trials should be evaluated
  Unexpected tissue responses
  Unexpected damage mechanisms
  Damage products in tissues
  F561
Further laboratory and animal testing needed?