## Slide 1: Testing for Irritation and Sensitization
Kenneth R. St. John, PhD
Chairman
F04.16 Subcommittee on Biocompatibility Test Methods

## Slide 2: Differences
Irritation usually reflects the tissue response at first exposure to the material or chemical
Sensitization usually refers to the response on repeat exposure, where the body has had the opportunity to develop an immune response
In humans, it is not always possible to know whether a patient has previously been exposed, so knowing the potential for sensitization may prevent mild or major complications

## Slide 3: Sensitizers in Medicine
Incomplete list
Nickel in metal alloys
Latex
  Gloves
  Vial seals
  Needles
  Adhesive bandages
Not necessarily addressing immune response to implants with skin tests, but the best options currently available

## Slide 4: Tests
Most experts and test documents suggest or require that materials or extracts be tested for irritation first at various doses and then the highest does for sensitization testing be the highest one that does not cause irritation
Many or most tests are for skin sensitization or irritation
Tests don’t really address implanted materials as well as maybe they could
Tests for immune response with implantation may be useful

## Slide 5: In Vivo Testing - Irritation
USP Irritation Test - Intracutaneous Injection Test
Mucous Membrane Irritation Test – No ASTM Standard
Edema and Erythema – Analogous to USP test but not stated to be an irritation test – toxicity is emphasized but scoring is identical to F719

## Slide 6: Human Testing?

## Slide 7: F719 - Irritation
Clip hair from back
  Option to add depilatory (non-Irritating)
Four sites, two intact and two abraded
Place samples, two control, two test
Cover with gauze and adhesive tape
Wrap trunk with polyethylene film
Remove after 24 hr
Score at 1 hr, 24 hr, 48 hr after removal

## Slide 8: F719
Calculate a primary irritation index, using formula in the specification, for each animal
Average for all animals

## Slide 9: Sensitization Testing
“Sensitization of hypersensitivity reactions usually occur as a result of prolonged contact with a chemical substance that interacts with the body’s immune system. Because most such reactions to biomaterials have been of the dermal cell-mediated type, rather than the humoral or antigen-antibody type, the skin of laboratory animals is used in sensitivity testing” – Richard F. Wallin (see reference at the end of talk)

## Slide 10: In Vivo Testing - Sensitization

## Slide 11: In Vivo Testing - Sensitization
“Biomaterials and other device materials are tested for the presence of sensitizing chemicals using guinea pigs. Guinea pig sensitization tests require six to eight weeks and thus take the longest time to complete of all the acute biocompatibility tests …” – Richard F. Wallin (see reference at the end of talk)

## Slide 12: F720 - Sensitization
Intradermal injection of an extract
  Three sites on each side of spine
  Adjuvant
  Sample without adjuvant
  Sample plus adjuvant
One week later
  Sample in petroleum jelly or liquid on filter paper applied at each site
  Leave for 48 hours
An additional week later
  Sample in petroleum jelly or liquid on filter paper applied at each site
  Leave for 24 hours
  Evaluate 1 Hr, 24 hr, 48 hr later
  Edema/Erythema scoring system (same as F719)

## Slide 13: F2147 - Sensitization
Intended to provide a different method of exposure that the Maximization test (F720), in which false positives were seen (both methods are internationally accepted)
A table of dose applications over about a 28 day period
Has a useful rationale and test comparison table in the appendix

## Slide 14: F2147 - Sensitization
Same scoring matrix as F719 and F720
More variations in the challenge types and intervals of repeated doses
Some experts consider this to be potentially more meaningful in actual application
All applications (with or without adjuvant) to the shaved skin but no skin injections

## Slide 15: F2148 - Sensitization
Lymph nodes in mice
Extract applied to both ears daily for three days
On day 6
  Prepare and inject a radiolabeled tracer (tritiated thymidine)
  Five hours later, euthanize mice
  Collect the auricular lymph node from each ear
  Mash or otherwise cut up lymph nodes into very small pieces
  Centrifuge and wash 3 times (PBS)
  Replace supernatant with Tricloroacetic acid, resuspend and precipitate overnight
  Replace TCA supernatant with fresh
  Resuspend and count redioactivity

## Slide 16: F2808 – Irritation Behind the Knee
Controversial – Not Accepted by US FDA
Possible that only the company that advocated for it actually uses it
Mostly used for over-the-counter OB/Gyn products
Places material behind the knee of human volunteers
  6 hours per day for 5 days
  Normal activities cause mechanical irritation in addition to testing chemical irritation

## Slide 17: Withdrawn Standards (2011)
Reflected an attempt to define in vitro tests for sensitization
Withdrawal was controversial – may need to be reconsidered in the future

## Slide 18: In Vitro Irritation and Sensitization
International focus to reduce the number of animals used in product testing
European Union Reference Laboratory for Alternatives to Animal Testing
  Develops and validates tests
  Has developed tests to assess irritation of cosmetics and chemicals
A US testing laboratory is working on validating and offering in vitro tests to the medical device industry

## Slide 19: In Vitro Irritation Testing
Reconstructed human epidermis (RhE)
Multi-Layered differentiated cell culture of human epidermis
Polar and non-polar extracts
MTT cytotoxicity test after exposure
  Assesses cell metabolic activity (viability)
May become a non-animal option for irritation testing
  Accepted in Europe for chemicals and cosmetics
  EU Test Method B.46 and OECD Test Guideline (TG) 439
Would likely become a standardized method if validated
Source: Medical Product Manufacturing News Sept/Oct 2015

## Slide 20: In Vitro Irritation Testing
Positive control apparently has been found – is an irritant for both polar and non-polar solvents
Validation will require proof of correlation with existing in vivo test results and/or clinical experience
At least one US contract testing laboratory is pursuing its acceptance

## Slide 21: In Vitro Sensitization Testing
Three different tests are being developed
One or a combination may end up being validated as non-animal tests
Completion and validation is still in progress

## Slide 22: Possible Reading
Rollins, T., “Better Animal Testing Alternatives Are Coming To US,” Medical Product Manufacturing News, p.18-19, September 2015
Lister, L., “Biocompatibility Testing – Tips for Avoiding Pitfalls – Part 2,” http://mddionline.com/print/6795
Wallin, R. F., “A Practical Guide to ISO 10993-10: Sensitization,” http://mddionline.com/print/2886
Wallin, R. F., “A Practical Guide to ISO 10993-10: Irritation,” http://mddionline.com/print/2906